## Slide 1
Index | General Requirements | Correspondence in the close-loop system | Priority
1 | No ventricular pace should happen during ventricle refractory | If State_RVA=ERP, no v_p should happen | 1
2 | Each atrial and ventricular event should be sensed by the corresponding lead | Interval between two Activation_RVA should between 500ms to 1000ms | 2
3 | Ventricular rate should be maintained between 60bpm and 120bpm | If Condition_NSR=true, Interval between two Activation_RVA shouldn’t be larger than Trest_SA; If Condition_Brady=true, Interval between two a_p should be equal to LRI timer of pacemaker | 3
4 | Without activity sensor, the pacemaker should not increase ventricular rate above it’s programmed LRI during Brady and above intrinsic heart rate during NSR | If Activation_SA=true, a_s=true; If Activation_RVA=true, v_s=true; | 4
5 | If the intrinsic heart rate is below some threshold, After each atrial event there should be a ventricular event within some interval(1:1 conduction) | If Interval between two Activation_SA is larger than 600, after each Activation_SA, there should be a Activation_RVA within [100, 150] | 5
6 | No activation conflict should happen within muscle tissue | musclepaths=path automata 1-10 in Fig 3(b); State_musclepaths should never be Double | 6